# Scope:

Glyphosate is best known as the active ingredient in Roundup-branded herbicides, and the herbicide used with “Roundup Ready” genetically modified organisms (GMOs) (Glyphosate Fact Sheet, 2020).

Due to overuse of the herbicide, glyphosate resistance has evolved in many weeds including several Amaranth species. Glyphosate resistance had been linked to an increase in the amount of target-site protein, 5-enolpyruvylshikimate-3-phosphate synthase (EPSPS). EPSPS’s protein levels and activity both increased as the number of EPSPS genomic copies increased (Gaines et al., 2010). The effect of the additional copies becomes additive and additional copies confer higher levels of resistance (Gaines et al., 2010).

NAGC validated a DNA-based test for determining EPSPS gene copy numbers using assays published by Gaines et al. (2010), Ma et al. (2013), and Giacomini et al. (2014). The test and estimated number of EPSPS gene copies for leaf samples will be normalized to that of a single copy genetic marker (acetolactate synthase; ALS) within the genome. This test will provide faster preliminary results to growers on the glyphosate resistance potential of pigweeds in their fields. Additionally, this test will provide weed scientists a rapid method to screen for herbicide resistance during their experiments in the greenhouse and field.

# Benefits to the Laboratory:

The quantitative PCR assay for measuring EPSPS gene amplification based on relative gene copy numbers will help assess the level of glyphosate resistance in amaranth populations.

# Definitions and Abbreviations:

ALS - acetolactate synthase

BLAST – Basic local alignment search tool

bp – base pair

Ct – cycling threshold

CPS - carbamoylphosphate synthetase

EPSPS – 5-enolpyruvylshikimate-3-phosphate synthase

IBFQ – Iowa black FQ, dark quencher from IDT

IBRQ – Iowa black RQ, dark quencher from IDT

IDT – Integrated DNA Technologies

ITS – internal transcribed spacer

MGW – molecular grade water

MU – Measurement of uncertainty

NA – nucleic acid

NAGC – National Agricultural Genotyping Center

NDSU – North Dakota State University

NP – non-probative

NT – not tested

NTC – no template control

PA – Palmer amaranth

PAC – positive amplification control

PCR – polymerase chain reaction

PO – Powell amaranth

qPCR – quantitative PCR

RC – reagent control

RR – redroot pigweed

RT – room temperature

SNP – single nucleotide polymorphisms

Und – undetermined cycling threshold value

WH - waterhemp

ZEN – internal quencher from IDT

# Equipment:

Allegra X-14 Centrifuge (Beckman Coulter)

Biomek® NXP Laboratory Automation Workstation (Beckman Coulter)

DNase/RNase-free pipette tips

DNase/RNase-free microcentrifuge tubes (2ml)

Freezer (-20ºC, -80ºC)

GenoGrinder

Heat Block (70°C, 80°C, 100C°; VWR)

MagnaRack Magnetic Separation Rack (ThermoFisher Scientific)

MiniSpin plus Centrifuge (Eppendorf)

PCR hood with UV light

Pipettes

QuantStudio 7 Flex Real-time PCR system with HRM software v1.2 (Life Technologies)

Refrigerator (4ºC)

Vortex

# Materials Used:

96-well Deep Plate (ThermoFisher Scientific)

Agarose

Aluminum sealing foil

CTAB lysis solution (BioChemica #A4150)

Ethanol

Ethidium bromide (10mg/ml solution)

GoTaq Probe qPCR Master Mix (Promega #A6102)

Isopropanol

Loading buffer

Maxwell® HT gDNA Blood System (Promega #A2671)

MicroAmp Fast 96-Well Reaction Plate (0.1ml, Applied Biosystems)

MicroAmp Optical Adhesive Film (Applied Biosystems)

Molecular grade water

TE buffer

Wizard® SV Gel and PCR Clean-Up System (Promega #A9281)

# SECTION 1: 	PRIMER INFORMATION AND PROBE DESIGN

The sequences of the primers for the EPSPS gene originated from the work of Gaines et al. (2010).  The EPSPS primer set, EPSF1 and EPSR8, are known to bind and amplify a 194bp product in multiple Amaranth species according to the literature and the cross-check of alignments using Geneious. The EPSPS primer set together along with primers that amplify a region of the plant ALS gene (ALSF2 and ALSR2) were used for quantitative PCR (Gaines et al., 2010; Giacomini et al. (2014). Because the ALS gene confers resistance to a variety of ALS-inhibiting herbicides, Ma et al. (2013) investigated an alternative reference gene, CPS that is not associated with herbicide resistance. The CPS gene encodes for the large subunit of carbamoylphosphate synthase. As explained below, the first trial of the CPS gene assay initially failed due to non-targeting primers (Section 3). However, once new genomes were available for bioinformatics (Montgomery et al. 2020), it became possible to develop new primers and probes to optimize the use of the CPS gene as the reference gene. Section 13 explains the comparative work between the ALS and CPS targets to support using CPS instead of ALS as the reference marker in the EPSPS-Copy Assay.

The assay performed by Gaines et al. (2010), as well as others, used SBYR green technology, which makes the assay require two separate reactions: one reaction for the EPSPS amplicon and a separate reaction for ALS or another single copy target. Instead of continuing with this 2-reaction protocol, we re-designed the assay and developed probes so that both markers could be co-amplified in the same reaction, resulting in easier assay setup and better design for high-throughput testing.

# SECTION 2: 	SYNTHETIC GBLOCK FORMULATION, RESUSPENSION, & DILUTIONS

Estimating copy numbers of the EPSPS gene in pigweeds requires a duplex assay that contains the internal control (ALS) to account for small differences in gDNA concentrations across leaf samples. Additionally, it is important to measure PCR efficiencies between the two assays as well as produce synthetic mixes that represent susceptibility (1:1 ratio of ALS to EPSPS). To this end, synthetic DNA (gBlocks) gene fragments for both EPSPS and ALS gene targets were synthesized by IDT. For each target, the sequence encompassed the primer binding locations as displayed by Geneious, adding an extra 25bp to each end. Equation 1 was used to obtain the starting concentration and copies/µl of the target fragments. The sequence, length, and specific accession number used are detailed in Appendix B.

Equation 1: Calculation of Copy Number from Molecular Weight:

(N) (6.022 x 1023 copies/mol) = Target Copies/µl

(MW) (1.0 x 109 ng/g) (V)

N = Amount of oligo in ng (provided by IDT on Specification sheet)

MW = Molecular weight of gBlock in g (provided by IDT on Specification sheet)

V = Volume used to resuspend oligo in µl

EPSPS

Upon arrival, the EPSPS gBlock (labelled as EPSP-Resistant) was resuspended in 50µl of TE-4 buffer with the calculated target copies/µl determined from the molecular weight and Avogadro’s number (copy number/mol) using Equation 1. Briefly, 50µl of TE-4 buffer was added to the gBlock tube and vortexed vigorously for 1 minute. The tube was briefly centrifuged to collect droplets to the bottom of the tube before incubating at 50°C for 20 minutes. Ten-fold serial dilutions were prepared in MGW containing 5% DMSO for use as positive amplification control material and stored at -20°C. This resulted in a more diluted stock than what was typically recommended from IDT.

ALS

Upon arrival, the ALS gBlock (labelled as syn-ALS) was resuspended in TE-4 buffer with the calculated target copies/µl determined from the molecular weight and Avogadro’s number (copy number/mol) using Equation 1. Briefly, 25µl of TE-4 buffer was added to the gBlock tube and vortexed vigorously for 1 minute. The tube was briefly centrifuged to collect droplets to the bottom of the tube before incubating at 50°C for 20 minutes. Ten-fold serial dilutions were prepared in MGW containing 5% DMSO for use as positive amplification control material and stored at -20°C.

Synthetic DNA EPSPS & ALS Working Stock Dilutions

Original stocks were diluted to reach 5ng/µl (EPSPS) and 10ng/µl (ALS), which equated to 1.99 x 1010 and 5.8 x 1010 copies/µl, respectively. These original stocks were further diluted in a 1:10 series to produce templates of 108 and 107 in 1mL volumes of 5% DMSO in MGW. Using the 107 stocks, a 1:10 dilution series of each synDNA were produced (106-101) in a final volume of 200µl of 5% DMSO in MGW. A 1:10 dilution series, beginning at 106 of a 1:1 mixture of both EPSPS and ALS synDNA, was created by adding 20µl of each 107 synDNA to 160µl of 5% DMSO in MGW. These synthetic stocks, particularly a mixture of both as well as diluted amaranth samples, were used to examine PCR efficiencies (see PCR Amplification Efficiency below).

# SECTION 3: 	EPSPS & INTERNAL CONTROL AMPLIFICATION TRIALS

### Part 1. Initial trial of the Primers and Probes

The primers and newly designed probes for detecting the EPSPS gene and reference genes of CPS and ALS were evaluated in the initial pilot assays (plates 051320_2RT-probe, 051320_2RTa, 051320_3RT-HRM, 063020_2RT-Probe, and 063020_4RT-HRM). The master mix reagents, concentrations, and standard cycling details are listed in Tables 3.1 and 3.2. The same formulation was used for either Melt Dr or probe-based reactions. Separate probes for each target were substituted within each singleplex reaction. MGW was substituted in place of probe volume.

| Table 3.1: The Glyphosate Resistance Assay master mix formulation – initial trial. |  |
| --- | --- |
| Master Mix Formulation | Reaction Volumes |
| Molecular Grade Water | 4.75µl |
| 2X Master Mix Buffer (Melt Dr or GoTaq Probe) | 12.5µl |
| 10µM Forward Primer [0.5µM] | 1.25µl |
| 10µM Reverse Primer [0.5µM] | 1.25µl |
| 10µM FAM Probe [0.1µM] | 0.25 µl |
| Template/Sample DNA | 5µl |
| Total Volume per Reaction | 25µl |

| Table 3.2: The ‘Herb Resist Assay’ thermal cycling program on the QuantStudio. |  |  |
| --- | --- | --- |
| Enzyme Activation | Denature | Anneal/Extend |
| HOLD | 40 CYCLES |  |
| 95°C for 2 min (probe) OR 95°C for 10 min (Melt Dr) | 95°C 30 sec | 60°C 1 min* |
| *Indicates where data signals were collected. |  |  |
| Standard heating properties were followed during the amplification phase, including the ramp rate of 1.6ºC/sec, unless otherwise indicated. |  |  |
| HRM melt curve analysis was separately after amplification for SYBR Green reactions. |  |  |

For the initial pilot assay, sample material tested included the EPSPS-R synDNA gBlock (concentration of 1x108 copies/µl) along with 10 DNA extracts from leaf samples that were submitted to NAGC for amaranth species identification. The amaranth species included 6 isolates of Palmer amaranth, 2 isolates of waterhemp, and a single isolate each of Powell amaranth and an unidentified pigweed species. The results of the EPSPS gene and reference gene detection assays are summarized in Table 3.3.

| Table 3.3: Initial pilot assay detection results for the EPSPS, ALS and CPS gene targets. |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Sample ID | Species ID | EPSPS Target |  | ALS Target |  | CPS Target |  |
|  |  | Probe Ct Values (051320_2RT) | SYBR Green Tm Values (051320_3RT) | Probe Ct Values (063020_2RT) | SYBR Green Tm Values (063020_4RT) | SYBR Green Ct Values (051320_2RTa) | SYBR Green Tm Values (051320_3RT) |
| EPSPS-R synDNA | EPSPS PAC | 8.393 | 81.639 | Und | Und | 32.949 | 79.097 |
| 19-0055-4 | Palmer | 16.718 | 81.267 | 23.957 | 81.614 | Und | 73.659 |
| 19-0068-1 | Waterhemp | 21.641 | 81.450 | 26.031 | 81.376 | 39.755 | 73.976 |
| 19-0078-2 | Waterhemp | 22.678 | 81.446 | 26.070 | 81.290 | Und | 65.284 |
| 19-0080-1 | Palmer | 18.423 | 81.259 | 23.846 | 81.498 | Und | 69.746 |
| 19-0085-3 | Powell | 23.918 | 81.501 | 24.937 | 81.314 | Und | 93.388 |
| 19-0093-3 | Palmer | 25.635 | 81.241 | 26.810 | 81.290 | Und | 65.262 |
| 19-0101-2 | Palmer | 23.000 | 81.245 | 23.882 | 81.424 | Und | 65.262 |
| 19-0104-4 | Palmer | 24.315 | 81.267 | 24.718 | 81.454 | Und | 74.628 |
| 19-0105-1 | Pigweed | 25.134 | 81.558 | 26.383 | 81.430 | Und | 65.205 |
| 19-0212-1 | Palmer | 25.907 | 81.291 | 27.229 | 81.457 | Und | 65.207 |
| NTC | N/A | Und | 78.356 | Und | Und | Und | 65.152 |

The EPSPS gene target assay produced Ct values for all samples at varying strengths ranging in Ct values from 16.718 to 25.907. The signal detected by the EPSPS-Pr probe appeared to be specific as indicated by the similar Tm profile as the positive control. All samples produced an EPSPS signal, indicating that all samples contained the EPSPS gene. According to Gaines et al. (2010), EPSPS is normally a low copy gene in plants, rice has one EPSPS locus, and Arabidopsis has two loci.

The CPS gene target assay did not appear to work. However, investigation later in the validation discovered that the purchased CPS-R primer was actually a reverse primer for the EPSPS target described within Ma et al. 2013. It seemed that during the initial purchase of the primers, the EPSPS reverse in the original study was purchased by NAGC and labeled CPS-R. As a result, ALS was used for the single copy reference gene. A supplemental validation will be required in order to use the CPS target as the housekeeping gene instead of ALS.

The ALS gene target assay also produced Ct values for all leaf samples, except the EPSPS-R synDNA. There was variation amongst the Ct values, ranging from 23.846 to 27.229. Since all samples were extracted using the same procedure (ChargeSwitch extraction kit), the variability could be from the differing concentrations of starting nucleic acid. This variability should be largely controlled during sample testing as long as the PCR efficiencies are relatively equal between the ALS and EPSPS targets. The analysis involves subtracting the ALS Ct values from the EPSPS Ct values, which normalizes the EPSPS copies within each sample.

Despite being used in previous studies, there were two main concerns of using the ALS region as the reference gene because of the selective pressure on this gene caused by ALS-inhibiting herbicide use. The first concern is the assay would be sub-optimal if it targets a region with high diversity or SNPs. According to the literature, there are 8 amino acid sites that when mutated lead to ALS herbicide inhibition (Murphy and Tranel 2019). None of the known mutation sites occur within the ALS amplicon. The second criticism is that there could be selection for copy number variation at the ALS gene, similar to the mechanism for glyphosate resistance and the EPSPS gene. Research so far has only indicated SNPs as the main target-site mechanisms for ALS-inhibitors. If ALS gene also has copy number variation this would result in an underestimation of EPSPS copies when both genes have elevated copies in the same plant. One could detect ALS copy number in instances where ALS gene had multiple copies and there was just one EPSPS gene copy within a pigweed. Under this scenario, the ALS target would have a lower Ct values than EPSPS target. The Ct values between the two targets will need to be carefully monitored during the validation, but as of writing and final approval of this validation, ALS copy number variation is only speculation.

### Part 2. EPSPS target primers and probe set

To be certain that the primer and probe set for EPSPS is detecting the EPSPS gene and not an Amaranth-specific gene, other non-Amaranth samples need to be tested, including those known to be glyphosate resistant and sensitive as well. As mentioned above, EPSPS is normally a low copy gene in plants, therefore any plant sample tested should provide a positive signal for EPSPS.

Using the same reaction and cycling conditions as stated in Tables 3.1 and 3.2, twenty-three miscellaneous DNA extracts from plants, insects, human, animal, and soils were screened for the EPSPS gene on plate 052620_2RT. The results are summarized in Table 3.4.

| Table 3.4: EPSPS detection in various DNA extracts. |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sample Name | DNA Type | EPSPS Result | Sample Name | DNA Type | EPSPS Result |
| Amaranth | Palmer seeds | Pos | Human | Human | Neg |
| Amaranth | Redroot seeds | Pos | Lentil | Plant | Neg |
| Barley | Plant | Neg | Pea | Plant | Neg |
| Canola | Plant | Neg | Rice | Plant | Neg |
| Chickpea | Plant | Neg | Soil-black | Soil | Neg |
| Corn | Plant | Neg | Soil-clay | Soil | Neg |
| (lot: 051622) |  |  | Soil-potting | Soil | Neg |
| Co-2 | Plant | Neg | Soil-sandy loam | Soil | Neg |
| 16-NP-341A2 | Plant | Neg | Salmon | Animal | Neg |
| 16-NP-342B2 | Plant | Neg | Soybean | Plant | Neg |
| 17-NP-9B | Plant | Neg | Spinach | Plant | Neg |
| 18-NP-282 | Plant | Neg | Sugarbeet | Plant | Neg |
| Honeybee | Insect | Neg | Wheat | Plant | Neg |

Other than the two Amaranth seed samples, none of the samples tested were positive for the EPSPS gene. This was initially concerning, since only Amaranth samples have been detected by the assay and the EPSPS gene is supposed to be a low copy gene in plants. It should be pointed out that rice is known to contain only a single copy of the EPSPS gene. Since a rice sample was tested and not detected, perhaps the EPSPS primers have variation among families. Performing a BLAST of the EPSPS amplified segment from Amaranth against soybean (Glycine max) found 3 accessions (XM_003516991, AK245902, and XM_003521809) that had percent identities of 77% with 87-88% query coverage. Geneious alignments of those same accessions against the EPSPS amplified region showed mismatches of 12, 8, and 6 nucleotides in the forward primer, probe and reverse primer, respectively. That number of mismatches in the primers and probe will likely prevent the amplification. Therefore, the assay in this validation appears to be specific to amplifying the EPSPS gene copies in pigweed species. The next step is to determine the optimal annealing temperature, primer, and probe concentrations for the duplex.

# SECTION 4: 	PCR OPTIMIZATION

Objective:

In contrast to most NAGC assays, the EPSPS-copy duplex requires a unique analysis that cannot favor amplification of one product over the other. If the duplex was optimized for just one marker (i.e., EPSPS or ALS), the results would be artificially skewed because the analysis takes the difference of Ct values between the two targets. Since a duplex assay is ideal, it makes sense to co-optimize the targets as a duplex rather than optimizing each singleplex, then duplex, as performed for other multiplex validations for pathogens. For example, it would not be appropriate to shift primer or probe concentrations of one marker over the other as it may artificially give amplification preference to one marker, resulting in the incorrect copy numbers for EPSPS of a sample with either inflated (or deflated) Ct values.

Procedure:

One of the hopes was to use the same thermal cycling conditions used for the Amaranth Species ID assay (Validation 18-007) so that species identification and glyphosate resistance could be determined on the same plate. Annealing temperature optimization assays were performed using the same master mix formulation and thermal cycling program modifying the annealing temperature from 58- 63°C. A combination of both real (pigweed) and synthetic (gBlock) samples were used for annealing temperature optimization. Samples were ran in duplicates or triplicate as well as 1:10 or 1:5 dilutions, depending on plate.

Results:

While there were obvious Ct differences between the three tested annealing temperatures for each assay when using the synthetic gBlocks (3-4 Cts), there were minimal differences in Ct values for the real waterhemp sample (19-0068-2.1), suggesting that 63°C should be further explored as the optimal annealing temperature for the glyphosate resistance assay (Table 4.1). Performing the PCR at 63°C would allow co-amplification with the Amaranth Species ID assay. The next step was to assess both gBlock mixtures and real sample dilutions to calculate PCR amplification efficiency for both targets (i.e., ALS and EPSPS) at 60°C (the published TA) and 63°C.

| Table 4.1. First trial of annealing temperature optimization with both synthetic and real samples. Plate runs are given below tested annealing temperatures. |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Sample Name | 58°C (081820_5RT) |  |  | 60°C (081720_2RT) |  |  | 63°C (081820_4RT) |  |  |
|  | EPSPS | ALS | ∆CT (A-E) | EPSPS | ALS | ∆CT (A-E) | EPSPS | ALS | ∆CT (A-E) |
| EPSPS_10_6 | 22.94 | Und | - | 19.02 | Und | - | 23.78 | Und | - |
| ALS_10_6 | Und | 18.09 | - | Und | 16.19 | - | Und | 18.13 | - |
| EPSPS-ALS-Mix-10_6 | 21.74 | 17.40 | -4.34 | 18.59 | 16.18 | -2.37 | 22.17 | 17.63 | -4.54 |
| 19-0068-2.1 (Wh) | 22.85 | 24.89 | 2.04 | 22.90 | 24.90 | 1.99 | 22.86 | 24.86 | 2.01 |

Trial 1 – 1:10 dilutions of an amaranth sample and a gBlock mixture at 60°C and 63°C

The initial two runs at 60°C and 63°C showed consistent patterns in a 1:10 dilution series of both synthetic gBlock mixes (106-101) and genomic DNA from a true waterhemp sample (neat – 10-5). However, the PCR amplification efficiencies for the assays were different and generally not ideal (< 95%) when calculated by either the gBlock or waterhemp sample (Table 4.2). When examining the data, most of the differences among replicates was found using gBlock synthetic DNA so the next trial will only use NAGC’s Maxwell extracts to further define PCR amplification efficiencies.

| Table 4.2. Dilution series for the initial examination of PCR amplification efficiencies at two annealing temperatures. Plate runs are given under the each tested annealing temperature. The samples that include EPSPS-ALS-Mix in their names are synthetic DNA mixtures, while the other samples listed are genomic DNA from previous cases or non-probative samples. |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Sample Name | 60°C (081920_3RT) |  |  | 63°C (081920_2RT) |  |  |
|  | EPSPS | ALS | ∆CT | EPSPS | ALS | ∆CT |
| EPSPS-ALS-Mix_10_6 (rep1) | 24.87 | 20.23 | -4.63 | 24.29 | 20.31 | -3.98 |
| EPSPS-ALS-Mix_10_6 (rep2) | 27.00 | 19.59 | -7.41 | 20.86 | 20.54 | -0.32 |
| EPSPS-ALS-Mix_10_5 (rep1) | 30.52 | 24.44 | -6.07 | 30.77 | 24.20 | -6.57 |
| EPSPS-ALS-Mix_10_5 (rep2) | 30.29 | 24.15 | -6.14 | 29.66 | 25.18 | -4.48 |
| EPSPS-ALS-Mix_10_4 (rep1) | 34.04 | 28.60 | -5.44 | 35.55 | 30.04 | -5.51 |
| EPSPS-ALS-Mix_10_4 (rep2) | 34.22 | 26.39 | -7.83 | 35.36 | 27.09 | -8.27 |
| EPSPS-ALS-Mix_10_3 (rep1) | Und | Und | - | Und | 30.58 | - |
| EPSPS-ALS-Mix_10_3 (rep2) | Und | 29.28 | - | 39.43 | Und | - |
| EPSPS-ALS-Mix_10_2 (rep1) | Und | Und | - | Und | Und | - |
| EPSPS-ALS-Mix_10_2 (rep2) | Und | Und | - | Und | Und | - |
| EPSPS-ALS-Mix_10_1 (rep1) | Und | Und | - | Und | Und | - |
| EPSPS-ALS-Mix_10_1 (rep2) | Und | Und | - | Und | Und | - |
| 19-0055-4.1 (Pa)_neat (rep1) | 17.65 | 23.81 | 6.16 | 17.14 | 23.83 | 6.69 |
| 19-0055-4.1 (Pa)_neat (rep2) | 17.62 | 23.77 | 6.15 | 17.15 | 23.87 | 6.72 |
| 19-0055-4.1 (Pa)_10_-1 (rep1) | 20.97 | 26.96 | 5.98 | 20.52 | 27.04 | 6.52 |
| 19-0055-4.1 (Pa)_10_-1 (rep2) | 20.97 | 27.05 | 6.09 | 20.51 | 27.02 | 6.51 |
| 19-0055-4.1 (Pa)_10_-2 (rep1) | 24.46 | 30.22 | 5.76 | 24.00 | 30.27 | 6.27 |
| 19-0055-4.1 (Pa)_10_-2 (rep2) | 24.45 | 30.30 | 5.86 | 23.98 | 30.38 | 6.40 |
| 19-0055-4.1 (Pa)_10_-3 (rep1) | 28.41 | 35.25 | 6.84 | 27.99 | 33.73 | 5.74 |
| 19-0055-4.1 (Pa)_10_-3 (rep2) | 28.48 | 34.17 | 5.69 | 27.93 | 34.57 | 6.64 |
| 19-0055-4.1 (Pa)_10_-4 (rep1) | 33.74 | 37.48 | 3.74 | 33.58 | 38.41 | 4.83 |
| 19-0055-4.1 (Pa)_10_-4 (rep2) | 33.84 | Und | - | 32.98 | Und | - |
| 19-0055-4.1 (Pa)_10_-5 (rep1) | 39.52 | Und | - | Und | Und | - |
| 19-0055-4.1 (Pa)_10_-5 (rep2) | 38.15 | Und | - | Und | Und | - |
| 20-NP-33o (RR) | 24.98 | 24.57 | -0.42 | 24.40 | 24.56 | 0.16 |
| 20-NP-36o (Wh) | 23.78 | 24.14 | 0.36 | 23.29 | 24.25 | 0.96 |

Trial 2 & 3 – 1:5 dilution of amaranth samples

The second run to monitor PCR amplification efficiencies used PCR conditions of Tables 4.1 and 4.2, which also included 9 reactions for the Amaranth Species ID assay. There were two changes to the second run to better assess PCR efficiencies. First, the dilution series was 1:5 dilutions (final volume 100µl) instead of 1:10 dilutions. Previously published research used a 1:5 dilutions series, described as 1x, 1/5x, 1/25x, and 1/125x (see Gaines et al. 2010). Second, three leaf samples were used that included Palmer amaranth (19-0055-4.1), a suspected redroot (20-NP-33o), and waterhemp (20-NP-39) to assess if there were any inter-species variation in the amplification efficiencies.

This third run used the same dilutions stocks as trial 2, but with reduced probe concentrations at 0.1µl instead of 0.2µM final conventions in the reactions. Reducing the probe concentrations had a no effects on the two amaranth samples (20-NP-33o & 20-NP-39) that contained low gene copies of EPSPS (Table 4.3). For the Palmer Amaranth samples, reducing the probe to 0.1µM had an effect on the EPSPS copies numbers presented for this sample. The EPSPS calculation is sensitive to slight differences in Ct values, especially as the Ct differences between ALS and EPSPS amplicons increase. However, this change from the estimated reduction from 64 to 32 copies (caused by the doubling effect) does not influence the overall qualitative reporting that this sample, 19-0055-4.1, is a glyphosate resistant plant.

| Table 4.3: Summary and comparison of final output (EPSPS copies) of the two plates with differing levels of probe concentrations. |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Probes at 0.2µM (082720_3RT) |  |  |  |  | Probes at 0.1µM (082720_4RT) |  |  |  |  |
| Sample Name | ALS Ct | EPSPS Ct | ∆Ct | Round ∆Ct | EPSPS copies (2∆Ct) | ALS Ct | EPSPS Ct | ∆Ct | Round ∆Ct | EPSPS copies (2∆Ct) |
| 19-0055-4.1 (Pa) 1/125 | 31.39 | 25.49 | 5.91 | 6 | 64 | 30.60 | 25.60 | 5.00 | 5 | 32 |
| 19-0055-4.1 (Pa) 1/125 | 31.11 | 25.46 | 5.64 | 6 | 64 | 30.46 | 25.58 | 4.87 | 5 | 32 |
| 19-0055-4.1 (Pa) 1/125 | 31.02 | 25.46 | 5.56 | 6 | 64 | 30.59 | 25.60 | 4.99 | 5 | 32 |
| 19-0055-4.1 (Pa) 1/25 | 28.61 | 22.78 | 5.83 | 6 | 64 | 28.09 | 22.81 | 5.28 | 5 | 32 |
| 19-0055-4.1 (Pa) 1/25 | 28.51 | 22.84 | 5.67 | 6 | 64 | 28.01 | 22.86 | 5.15 | 5 | 32 |
| 19-0055-4.1 (Pa) 1/25 | 28.52 | 22.80 | 5.72 | 6 | 64 | 28.08 | 22.81 | 5.27 | 5 | 32 |
| 19-0055-4.1 (Pa) 1/5 | 26.21 | 20.41 | 5.80 | 6 | 64 | 25.73 | 20.44 | 5.29 | 5 | 32 |
| 19-0055-4.1 (Pa) 1/5 | 26.21 | 20.42 | 5.79 | 6 | 64 | 25.71 | 20.41 | 5.31 | 5 | 32 |
| 19-0055-4.1 (Pa) 1/5 | 26.12 | 20.36 | 5.76 | 6 | 64 | 25.64 | 20.38 | 5.26 | 5 | 32 |
| 19-0055-4.1 (Pa) Neat | 23.88 | 17.95 | 5.93 | 6 | 64 | 23.30 | 17.92 | 5.39 | 5 | 32 |
| 19-0055-4.1 (Pa) Neat | 23.90 | 18.03 | 5.86 | 6 | 64 | 23.36 | 18.05 | 5.31 | 5 | 32 |
| 19-0055-4.1 (Pa) Neat | 23.92 | 18.04 | 5.88 | 6 | 64 | 23.34 | 18.05 | 5.29 | 5 | 32 |
| 20-NP-33o 1/125 | 31.69 | 32.53 | 0.84 | 1 | 2 | 31.69 | 32.75 | 1.05 | 1 | 2 |
| 20-NP-33o 1/125 | 31.98 | 32.62 | 0.64 | 1 | 2 | 31.63 | 32.63 | 1.00 | 1 | 2 |
| 20-NP-33o 1/125 | 31.99 | 32.45 | 0.45 | 1 | 2 | 31.46 | 32.62 | 1.16 | 1 | 2 |
| 20-NP-33o 1/25 | 29.48 | 30.14 | 0.66 | 1 | 2 | 29.27 | 30.21 | 0.95 | 1 | 2 |
| 20-NP-33o 1/25 | 29.47 | 30.14 | 0.67 | 1 | 2 | 29.17 | 30.15 | 0.98 | 1 | 2 |
| 20-NP-33o 1/25 | 29.41 | 30.07 | 0.66 | 1 | 2 | 29.04 | 30.02 | 0.98 | 1 | 2 |
| 20-NP-33o 1/5 | 27.03 | 27.70 | 0.67 | 1 | 2 | 26.78 | 27.72 | 0.94 | 1 | 2 |
| 20-NP-33o 1/5 | 27.06 | 27.75 | 0.70 | 1 | 2 | 26.79 | 27.77 | 0.97 | 1 | 2 |
| 20-NP-33o 1/5 | 27.02 | 27.81 | 0.78 | 1 | 2 | 26.81 | 27.76 | 0.95 | 1 | 2 |
| 20-NP-33o Neat | 24.72 | 25.39 | 0.67 | 1 | 2 | 24.33 | 25.29 | 0.96 | 1 | 2 |
| 20-NP-33o Neat | 24.68 | 25.35 | 0.67 | 1 | 2 | 24.31 | 25.31 | 1.00 | 1 | 2 |
| 20-NP-33o Neat | 24.68 | 25.37 | 0.69 | 1 | 2 | 24.32 | 25.31 | 1.00 | 1 | 2 |
| 20-NP-39 1/125 | 31.44 | 29.44 | 2.00 | 2 | 4 | 31.07 | 29.66 | 1.41 | 1 | 2 |
| 20-NP-39 1/125 | 31.70 | 29.65 | 2.05 | 2 | 4 | 31.32 | 29.67 | 1.64 | 2 | 4 |
| 20-NP-39 1/125 | 31.54 | 29.64 | 1.90 | 2 | 4 | 31.28 | 29.71 | 1.57 | 2 | 4 |
| 20-NP-39 1/25 | 29.11 | 27.19 | 1.92 | 2 | 4 | 28.79 | 27.21 | 1.59 | 2 | 4 |
| 20-NP-39 1/25 | 29.13 | 27.26 | 1.88 | 2 | 4 | 28.80 | 27.21 | 1.59 | 2 | 4 |
| 20-NP-39 1/25 | 29.04 | 27.15 | 1.90 | 2 | 4 | 28.78 | 27.23 | 1.54 | 2 | 4 |
| 20-NP-39 1/5 | 26.83 | 24.85 | 1.98 | 2 | 4 | 26.56 | 24.90 | 1.66 | 2 | 4 |
| 20-NP-39 1/5 | 26.85 | 24.84 | 2.01 | 2 | 4 | 26.49 | 24.88 | 1.61 | 2 | 4 |
| 20-NP-39 1/5 | 26.87 | 24.88 | 1.99 | 2 | 4 | 26.59 | 24.93 | 1.65 | 2 | 4 |
| 20-NP-39 Neat | 24.58 | 22.47 | 2.11 | 2 | 4 | 24.21 | 22.46 | 1.75 | 2 | 4 |
| 20-NP-39 Neat | 24.53 | 22.47 | 2.06 | 2 | 4 | 24.21 | 22.50 | 1.71 | 2 | 4 |
| 20-NP-39 Neat | 24.54 | 22.52 | 2.03 | 2 | 4 | 24.20 | 22.53 | 1.68 | 2 | 4 |

Conclusions:

The two assays (ALS and EPSPS) should maintain similar reagent concentrations for estimating EPSPS gene copies across samples. Therefore, only annealing temperature and probe concentrations were examined for optimization purposes. Additionally, this will allow both the Amaranth Species ID and Glyphosate Resistance assay to be co-amplified on the same run (see co-amplification on plate 082720_3RT). Also, the qualitative results that are reported to clients does not change by reducing the probe concentration. As a final confirmation to these PCR conditions, PCR amplification efficiencies should be calculated using the same data above, but analyzed as a standard curve. The fact that all four 1:5 dilutions for the three leaf samples report identical EPSPS gene copies suggests that PCR amplification efficiencies are suitable under the recommended PCR conditions.

# SECTION 5: 	PCR AMPLIFICATION EFFICIENCY OF EPSPS-Copy ASSAY

Objective:

To properly estimate EPSPS copies and subsequent glyphosate resistance in pigweeds, both targets within the multiplex must have relatively similar PCR amplification efficiencies.

Procedure:

The previously amplified 1:10 and 1:5 dilutions of the synthetic gBlocks and amaranth leaf samples were used to examine individual and collective amplification efficiencies using the Standard Curve methods within the QuantStudio Real-Time PCR software. PCR amplification efficiencies are calculated by the following equation: E = -1+10(-1/slope).

Results:

Overall, the 1:5 dilution series of leaf material provided the best measures of PCR amplification efficiencies for EPSPS (Table 5.1) and ALS (Table 5.2). There was a 2% drop in PCR efficiency values calculated for the EPSPS assay when probe was reduced to 0.1µM in the reactions. This calculation is very sensitive due to the small changes in Ct values when reducing the probe concentrations. Despite this drop to a 92% efficiency, it is within the recommended range for qPCR analyses (>90%). For ALS, efficiency was maintained at 96% despite the drop in probe concentrations, indicating the internal control as a robust assay for the glyphosate resistance. Ideally, the efficiencies would be similar between the two targets, but a range of 3-4% was also found by Gaines et al. (2010). Efficiencies also improved for 19-0055-4.1 when 1:5 dilutions were used instead of 1:10 for the calculations, perhaps due to the reduction of pipetting error with greater sample volume added to diluent for 1:5 dilutions. Synthetic DNA had the lowest PCR efficiencies for the samples serially diluted, which may be an intrinsic quality of gBlocks. Thus, it is recommended to use leaf samples for QC samples over gBlocks, especially if PCR efficiencies must be calculated.

The low R2 values (0.45 and 0.46; Table 5.1) were expected in the combined leaf datasets because the samples have different EPSPS copies per genome. For instance, 19-005-4.1 has a greater number of EPSPS gene copies (~60) compared to 20-NP-39 (~2), but both were included in the single analysis. Due to these differences in EPSPS gene copies, neat samples contain different starting amounts of the EPSPS target (and subsequent Ct values), resulting in larger distances between the data points at each dilution along the regression line in the 1:10 and 1:5 dilution series. This scatter in the data reduces R2 values in the combined leaf standard curve analyses. For leaf samples with varying copies of EPSPS, it is more appropriate to examine standard curves for each leaf sample separately, which have the expected R2 at or nearing 1, regardless of EPSPS copy number (see Table 5.1). Note that the combined ALS R2 values were also lower (0.91-0.97; Table 5.2), albeit not as extreme as the EPSPS, because leaf samples contain only a single copy of ALS. The lower R2 values for ALS are likely due to the small differences in the starting gDNA concentrations in each leaf extract that arise from the independent DNA extractions performed on each sample.

| Table 5.1. Summary of PCR amplification efficiencies for the EPSPS target from QuantStudio software. Values include those from 1:10 dilutions series (plate: 081920_2RT) of leaf and synthetic DNA and 1:5 dilutions of leaf DNA that had two different probe concentrations of 0.2µM (plate: 082720_3RT) and 0.1µM (plate: 082720_4T). |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Dilution Series | Sample | EPSPS (0.2µM Probes) |  |  | EPSPS (0.1µM Probes) |  |  |
|  |  | Slope | Efficiency | R2 | Slope | Efficiency | R2 |
| 1:5 | 19-0055-4.1 (Pa) | -3.54 | 91.50 | 0.99 | -3.61 | 89.40 | 0.99 |
|  | 20-NP-39 (Wh) | -3.38 | 97.70 | 1.00 | -3.41 | 96.30 | 1.00 |
|  | 20-NP-33 (RR) | -3.42 | 96.30 | 1.00 | -3.50 | 93.10 | 0.99 |
|  | Combined | -3.45 | 94.90 | 0.45 | -3.51 | 92.85 | 0.46 |
| 1:10 | 19-0055-4.1 (Pa) | -3.97 | 78.50 | 0.99 | - | - | - |
|  | EPSPS-ALS 10_6 Mix | -5.77 | 48.95 | 0.95 | - | - | - |
|  | Combined | -3.52 | 92.30 | 0.49 | - | - | - |

| Table 5.2. Summary of PCR amplification efficiencies for the ALS target. Values include those from 1:10 dilutions series (plate: 081920_2RT) of leaf and synthetic DNA and 1:5 dilutions of leaf DNA that had two different probe concentrations of 0.2µM (plate: 082720_3RT) and 0.1µM (plate: 082720_4T). |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Dilution Series | Sample | ALS (0.2µM Probes) |  |  | ALS (0.1µM Probes) |  |  |
|  |  | Slope | Efficiency | R2 | Slope | Efficiency | R2 |
| 1:5 | 19-0055-4.1 (Pa) | -3.43 | 95.80 | 0.99 | -3.47 | 94.00 | 0.99 |
|  | 20-NP-39 (Wh) | -3.33 | 99.60 | 0.99 | -3.33 | 99.70 | 0.99 |
|  | 20-NP-33 (RR) | -3.46 | 94.70 | 0.99 | -3.43 | 95.60 | 0.99 |
|  | Combined | -3.41 | 96.46 | 0.97 | -3.41 | 96.53 | 0.98 |
| 1:10 | 19-0055-4.1 (Pa) | -3.56 | 90.72 | 0.99 | - | - | - |
|  | EPSPS-ALS 10_6 Mix | -3.58 | 90.20 | 0.93 | - | - | - |
|  | Combined | -3.79 | 83.50 | 0.91 | - | - | - |

Conclusions:

PCR efficiencies for both ALS and EPSPS assays are within the appropriate range for gene copy estimates to be calculated.

# SECTION 6: 	SPECIFICITY ANALYSIS

Objective:

The purpose of the specificity analysis was to evaluate the primers and probes, checking for cross-reaction to non-specific targets. Sound primer and probe design for the target of interest in the assay will ensure minimal cross-reactivity from non-specific sources, many of which can be found ubiquitous in nature.

Procedure:

DNA from laboratory, agriculturally, and environmentally-related samples were previously extracted using established laboratory protocols. These samples were screened using the EPSPS-Copy assay to verify cross-reactivity with other DNA samples that occur in the laboratory.

Results & Conclusions:

The specificity analysis revealed one potential non-target amplification for one of the gene targets, EPSPS (Table 6.1). No other plant, besides pigweed species produced both ALS and EPSPS signals. However, the template containing Fusarium graminerium produced a high Ct value of 38 for EPSPS. For the first test, there was no amplification of the ALS target, which would require retesting if the sample was an unknown. The source of the EPSPS amplification is unknown and could be cross-reaction with F. graminerium or some plant-based material that the pathogen was cultured on. A Ct of 38 would be suspect even from a plant sample provided that NAGC’s archived samples of leaf material produce Cts within the 20s. A second run (021221_2RT) of the F. graminerium showed amplification of both targets (EPSPS = 37 and ALS = 38), suggesting potential contamination of this sample with Amaranth DNA or cross-reaction with this extract. It seems unlikely that both targets would cross-react within a fungal species, especially since the other Fusarium species tested did not show weak amplification. A new lot of Fusarium graminerium (agar plugs stored at -80°C) was tested using the EPSPS-Copy assay. This independent extract also had a weak positive signal for ALS (Ct = 38), indicating persistent cross reactivity (32221_2RT). Acetolactate synthase (ALS) is known to exist in plants, bacteria, and fungi. Thus, it is possible that the ALS assay is able to cross react with fungal species. There is little concern of this cross reaction causing issues at NAGC because the EPSPS amplicon is missing in this samples, which is needed for full analysis.

| Table 6.1 Specificity results (Ct values) for ALS and EPSPS targets. Green shading indicates the potential cross-reaction with a Fusarium species. Data analyzed from plate 090220_1RT. |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Sample Name | ALS | EPSPS |  | Sample Name | ALS | EPSPS |
| Palmer amaranth | 24.117 | 21.101 |  | Human DNA | Und | Und |
| Amaranth PAC | 25.041 | 22.341 |  | Soybean | Und | Und |
| 20-NP-39 | 24.466 | 22.642 |  | LAB Ctrl | Und | Und |
| Redroot pigweed | 24.796 | 25.679 |  | DCV PAC | Und | Und |
| F. graminerium | Und | 38.434 |  | Neg Bee (18-NP-289) | Und | Und |
| ABL2 PAC | Und | Und |  | Sugarbeet | Und | Und |
| GLX PAC | Und | Und |  | Corn (CO-2) | Und | Und |
| Potting Soil (16-NP-318) | Und | Und |  | NCLB PAC | Und | Und |
| NTC | Und | Und |  | Barley | Und | Und |
| CLRDV PAC | Und | Und |  | Lentil | Und | Und |
| AES PAC | Und | Und |  | Wheat DNA | Und | Und |
| F. subglutinans | Und | Und |  | 19-SCN-19 | Und | Und |
| Salmon sperm DNA | Und | Und |  | PS PAC | Und | Und |
| Sandy Loam (16-NP-108) | Und | Und |  | Canola | Und | Und |
| DIP PAC | Und | Und |  | Pea | Und | Und |
| KNN PAC | Und | Und |  | Elk Muscle | Und | Und |
| F. verticillioides | Und | Und |  | 19-SNC-4 | Und | Und |
| Spinach | Und | Und |  | Puccinia PAC | Und | Und |
| GW PAC | Und | Und |  | Chickpea | Und | Und |
| ISL1 PAC | Und | Und |  | Rice | Und | Und |
| Clay Soil (16-NP-317) | Und | Und |  |  |  |  |

# SECTION 7: 	SENSITIVITY & MIXTURE ANALYSES

Objective:

After confirming the primer and probe sets were specific for their intended target, the next step in the validation was to assess their sensitivity. The sensitivity study was designed to evaluate the amplification of the targets over a range of template concentrations. Since the EPSPS-Copy assay relies on the Ct difference between EPSPS and ALS amplification, sensitivity of the assay should incorporate differential dilutions of the two targets within a sample to investigate how much of an amplification gap (i.e., Ct difference) between the targets can be reliably detected using standard PCR conditions.

Procedures:

Several runs were performed by two different analysts to investigate the sensitivity of this duplex assay. One of the main goals was to determine how much of a Ct difference between the amplicons was detectable using the optimized conditions above. Both of these plate runs used EPSPS and ALS synthetic DNA stocks to make the mixtures of varying dilutions. Within the first trial run, a confirmed waterhemp sample (20-NP-39) was used to examine the overall sensitivity of the EPSPS-Copy assay by performing 1:10 dilutions (Neat – 10-7). Additional synDNA samples of the single and mixed targets were also tested for sensitivity analyses.

Trial 1 – Previous working dilutions of EPSPS and ALS synDNAs were used in mixtures with a final volume of 100µl MGW. The working stock concentrations as well as the final dilutions for the mixture samples are provided in Table 7.1

| Table 7.1: Sample design for the mixtures produced for Trial 1. |  |  |  |
| --- | --- | --- | --- |
| Sample (EPSPS:ALS) | EPSPS stock concentration | ALS stock concentration | Final Mixtures |
| 1 (10:1) | 1.99 × 107 | 5.81 × 106 | EPSPS 106 ALS 105 |
| 2 (100:1) | 1.99 × 107 | 5.81 × 105 | EPSPS 106 ALS 104 |
| 3 (1000:1) | 1.99 × 107 | 5.81 × 104 | EPSPS 106 ALS 103 |
| 4 (10000:1) | 1.99 × 107 | 5.81 × 103 | EPSPS 106 ALS 102 |
| 5 (1:1) | 1.99 × 107 | 5.81 × 107 | EPSPS 106 ALS 106 |

Trial 2 – A fresh batch of 1:10 dilutions were made of each synDNA stock (107-103) in 100µl MGW containing 5% DMSO. Along with diluting ALS synDNA to create a differential mixture of the two targets, a set of mixture samples were created that kept ALS concentration constant while reducing EPSPS concentrations. These dilutions were then used for the mixture studies as described in Table 7.2.

| Table 7.2: Sample design for the mixtures produced for Trial 2. |  |  |  |
| --- | --- | --- | --- |
| Sample (EPSPS:ALS) | EPSPS stock concentration | ALS stock concentration | Final Mixtures |
| 1 (10:1) | 1.99 × 107 | 5.81 × 106 | EPSPS 106 ALS 105 |
| 2 (100:1) | 1.99 × 107 | 5.81 × 105 | EPSPS 106 ALS 104 |
| 3 (1000:1) | 1.99 × 107 | 5.81 × 104 | EPSPS 106 ALS 103 |
| 4 (1:10) | 1.99 × 106 | 5.81 × 107 | EPSPS 105 ALS 106 |
| 5 (1:100) | 1.99 × 105 | 5.81 × 107 | EPSPS 104 ALS 106 |
| 6 (1:1000) | 1.99 × 104 | 5.81 × 107 | EPSPS 103 ALS 106 |

Results & Conclusions:

The 1:10 dilutions of the waterhemp sample provided consistent EPSPS gene copies results through three serial dilutions of the original sample (Table 7.3). These results indicated that the protocols and extraction methods for the Amaranth Species ID test (SOP 5.11) are also ideal for the EPSPS-Copy assay. Additionally, the consistency of the reported EPSPS gene copies (3.7 - 4.8) across the dilution series suggested no impact of PCR inhibitors within the extraction method and neat template can be used for this assay.

| Table 7.3 Sensitivity results from the 1:10 dilution series of a single waterhemp sample (20-NP-39). Provided are the Ct values for ALS and EPSPS targets, the differences of these two values (∆CT), the absolute difference, and the raw estimates of EPSPS gene copies (2∆CT). The green shaded examples show the limit of detection for both targets. Data from plate 090420_3RT. |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sample Name | ALS (Ct) | EPSPS (Ct) | ∆CT (intra-assay) | Abs (∆CT) | EPSPS copies (2∆CT) |
| 20-NP-39 Neat | 24.273 | 22.232 | -2.041 | 2.0411 | 4.115 |
| 20-NP-39 Neat | 24.151 | 22.235 | -1.915 | 1.9153 | 3.772 |
| 20-NP-39_10-1 | 27.563 | 25.670 | -1.894 | 1.8936 | 3.716 |
| 20-NP-39_10-1 | 27.547 | 25.610 | -1.937 | 1.9365 | 3.828 |
| 20-NP-39_10-2 | 31.300 | 29.192 | -2.108 | 2.1077 | 4.310 |
| 20-NP-39_10-2 | 31.061 | 29.206 | -1.855 | 1.855 | 3.618 |
| 20-NP-39_10-3 | 34.360 | 32.468 | -1.892 | 1.8916 | 3.710 |
| 20-NP-39_10-3 | 34.928 | 32.646 | -2.282 | 2.2822 | 4.864 |
| 20-NP-39_10-4 | Und | 36.476 | - | - | - |
| 20-NP-39_10-4 | 38.236 | 36.869 | -1.367 | 1.3671 | 2.580 |
| 20-NP-39_10-5 | Und | 39.180 | - | - | - |
| 20-NP-39_10-5 | Und | Und | - | - | - |
| 20-NP-39_10-6 | Und | Und | - | - | - |
| 20-NP-39_10-6 | Und | Und | - | - | - |
| 20-NP-39_10-7 | Und | Und | - | - | - |
| 20-NP-39_10-7 | Und | Und | - | - | - |

In contrast, the synthetic DNA, specifically EPSPS, did not show the 3 Ct increase for the dilutions (Table 7.4). The ALS synDNA had the expected pattern of Ct increases and both replicates amplified down to 104. These dilutions were the stocks used to make the artificial samples with elevated gene copies in the Trial 2 mixture study. Therefore, the resulting mixture studies may not display the spread of Cts that are expected from the dilution measures. Nevertheless, the mixture trials below help to examine how much spread of Cts between the two targets can be measured with the EPSPS-Copy assay.

| Table 7.4 Sensitivity results from the 1:10 dilution series of EPSPS and ALS synDNA targets. The green shaded examples show the limit of detection where both replicates provided Ct values. Data from single target templates on 112520_1RT. |  |  |  |
| --- | --- | --- | --- |
| ALS Copy Number | ALS (Ct) singleplex | EPSPS Copy Number | EPSPS (Ct) singleplex |
| 5.81 × 107 | 26.939 | 1.99 × 107 | 25.914 |
| 5.81 × 107 | 26.916 | 1.99 × 107 | 26.833 |
| 5.81 × 106 | 28.583 | 1.99 × 106 | 32.146 |
| 5.81 × 106 | 28.519 | 1.99 × 106 | 26.444 |
| 5.81 × 105 | 32.786 | 1.99 × 105 | 39.231 |
| 5.81 × 105 | 32.336 | 1.99 × 105 | Und |
| 5.81 × 104 | 35.247 | 1.99 × 104 | Und |
| 5.81 × 104 | 34.616 | 1.99 × 104 | Und |
| 5.81 × 103 | 38.327 | 1.99 × 103 | Und |
| 5.81 × 103 | Und | 1.99 × 103 | Und |

Within the mixed synDNA samples, a Ct differential of 10 was possible, which equated to 1,265 copies of EPSPS gene (2∆CT method; Table 7.5). This value was arguably inflated beyond the biological possibility for multiple gene copies occurring in nature. Nevertheless, this provides confidence that the assays can detect large Ct variation caused by large differences in the relative gene copies of EPSPS with respect to ALS, a single-copy reference gene.

In Trial 1, the EPSPS target was maintained at a constant concentration (1.99 × 106 copies/µl) as ALS was reduced 10-fold. Despite being held at a constant concentration, there was still variation in the EPSPS Ct values (range: 24-27). Likewise, Trial 2 displayed Ct variation for the target held at a constant concentration by 2 Cts. This may be due to pipetting error because each mixture (e.g., 1EPSPS:1ALS, 1EPSPS:10ALS, etc) were prepared individually. The ALS target showed the expected increase in Ct values with each 10-fold drop of synDNA in Trial 2, but was more sporadic in Trial 1. Despite these discrepancies within the mixture study, the trials using synthetic DNA provided evidence well beyond the biological expectations (i.e., 1000 gene copies) that high EPSPS gene copies can be detected using this probe-based assay.

| Table 7.5 Sensitivity results from the two trials on the mixture analyses. Provided are the Ct values for ALS and EPSPS targets, the differences of these two values (∆CT), the absolute difference, and the raw estimates of ALS or EPSPS gene copies (2∆CT). The positive values within the ∆CT column indicate where there was more ALS than EPSPS product (i.e., low Ct value). In these instances the 2∆CT reflects the levels of ALS gene copies, which is not expected in real samples because ALS should be a single copy gene and should not have more copies than EPSPS. Data from plates 090420_3RT (Trial 1) and 112520_1RT (Trial 2). |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Trial | Sample Name | ALS (Ct) | EPSPS (Ct) | ∆CT (intra-assay) | Abs (∆CT) | EPSPS (ALS) copies (2∆CT) |
| 1 | EPSPS:ALS_1:1 | 23.006 | 24.952 | 1.945 | 1.945 | 3.851 (ALS) |
|  | EPSPS:ALS_1:1 | 24.196 | 25.075 | 0.879 | 0.879 | 1.840 (ALS) |
|  | EPSPS:ALS_10:1 | 34.108 | 26.894 | -7.214 | 7.214 | 148.501 |
|  | EPSPS:ALS_10:1 | 35.246 | 24.940 | -10.306 | 10.306 | 1,265.845 |
|  | EPSPS:ALS_10:1 | 34.768 | 26.167 | -8.601 | 8.601 | 388.333 |
|  | EPSPS:ALS_10:1 | 34.559 | 27.295 | -7.264 | 7.264 | 153.737 |
|  | EPSPS:ALS_100:1 | Und | 27.366 | - | - | - |
|  | EPSPS:ALS_100:1 | Und | 27.585 | - | - | - |
|  | EPSPS:ALS_1000:1 | 32.413 | 26.978 | -5.435 | 5.435 | 43.262 |
|  | EPSPS:ALS_1000:1 | 32.312 | 27.681 | -4.631 | 4.630 | 24.775 |
|  | EPSPS:ALS_10000:1 | Und | 26.752 | - | - | - |
|  | EPSPS:ALS_10000:1 | Und | 27.382 | - | - | - |
| 2 | EPSPS:ALS_1000:1 | 38.442 | 30.524 | -7.918 | 7.918 | 241.906 |
|  | EPSPS:ALS_1000:1 | 36.066 | 32.969 | -3.097 | 3.097 | 8.557 |
|  | EPSPS:ALS_100:1 | 34.941 | 32.859 | -2.082 | 2.082 | 4.234 |
|  | EPSPS:ALS_100:1 | 35.178 | 31.295 | - 3.882 | 3.882 | 14.747 |
|  | EPSPS:ALS_10:1 | 32.802 | 29.969 | 2.833 | 2.832 | 7.124 (ALS) |
|  | EPSPS:ALS_10:1 | 33.081 | 37.452 | -4.371 | 4.370 | 20.685 |
|  | EPSPS:ALS_1:1 | 28.805 | 31.979 | 3.174 | 3.173 | 9.023 (ALS) |
|  | EPSPS:ALS_1:1 | 28.778 | 31.303 | 2.525 | 2.524 | 5.754 (ALS) |
|  | EPSPS:ALS_1:10 | 28.819 | 35.262 | 6.443 | 6.442 | 86.987 (ALS) |
|  | EPSPS:ALS_1:10 | 28.759 | 36.333 | 7.574 | 7.573 | 190.535 (ALS) |
|  | EPSPS:ALS_1:100 | 29.007 | 38.579 | 9.572 | 9.571 | 761.047 (ALS) |
|  | EPSPS:ALS_1:100 | 29.325 | Und | - | - | - |
|  | EPSPS:ALS_1:1000 | 29.531 | Und | - | - | - |
|  | EPSPS:ALS_1:1000 | 29.361 | Und | - | - | - |

# SECTION 8: 	EPSPS COPY NUMBER VARIATION IN OLD VS NEW LEAVES FROM SAME PLANT

Objective:

A recent study of Giacomini et al. (2019) examined the heritability and segregation patterns of EPSPS copy number within Palmer amaranth. This study found that EPSPS copy number variation did not follow an expected Mendelian single-gene inheritance pattern and that cloned cells taken from different parts of an individual plant showed varying EPSPS copy numbers. These results suggest that EPSPS amplification within a plant may be variable, localized, and random within a plant. The working hypotheses based on Koo et al. (2018)’s cellular analyses using FISH (fluorescence in situ hybridizations) assays, is that EPSPS copies are located within extra-chromosomal circular DNA (eccDNA) molecules in Palmer amaranth plants. The eccDNA is tethered to the primary chromosomes, allowing it to be passed through mitotic as well as meiotic cell division. From a testing or diagnostic testing standpoint, this could be troublesome because it could lead to false negative diagnosis for glyphosate susceptibility if the portion of a resistant plant sampled had low EPSPS copies. Additionally, plants may have the ability to boost division of specific cells containing these eccDNA when exposed to glyphosate, resulting in a rapid change within a single generation. One suggestion was to test multiple leaves along the meristem (i.e., stalk) as their examination of 6 different leaves on the vertical meristem from one plant ranged in EPSPS copies from 2-32. This range of EPSPS copies equates to 1-5 Ct differences between the ALS and EPSPS targets.

Procedures:

To explore some of these concerns independently, several fresh amaranth samples were collected from a soybean field near Kindred, ND. This soybean field was known to be sprayed with glyphosate (round up) and samples were collected late in the season (August 2020) that were likely herbicide escapees (Figure 8.1). From these four plants, an old and young leaf were collected and a standard size hole punch size sample was used for extractions and PCR analysis. One of the extracts from each plant was tested using the Amaranth Species ID assays (SOP 5.11) to identify the pigweed species.

Figure 8.1. Soybean field and pigweed escapees that were collected for this experiment. The red circle in the right image shows evidence that the pigweed (waterhemp) was exposed to a stressor during the growing season, which included applications of roundup (glyphosate).

Results & Conclusions:

There was no obvious difference between sampling an old or young leave from the pigweed escapees as all samples estimated 1-2 EPSPS copies per sample (Table 8.1). Unfortunately, none of the samples had elevated EPSPS gene copies so it is unknown if there are more variation of EPSPS copies across leaves in a plant that was reported with elevated EPSPS copies. Besides the potential for EPSPS copy number variation within a plant, there are several reasons why all of the tested pigweeds that appeared to escape herbicide applications did not show elevated EPSPS. The application may not have been effective due to inclement weather, blocked sprayer, and poor timing of herbicide. This might explain the stressed meristems in Fig. 8.1. Additionally, there are other known target-site and non-target site mechanisms for glyphosate resistance, which may explain the resistance of these plants. A greenhouse bioassay and the EPSPS-Copy assay will examine a subset of seeds collected and grown from sunflower screenings to better describe the frequency of the EPSPS target-site mechanism (i.e., copy number variation).

| Table 8.1. Results from examining EPSPS copies in old (o) and young (y) leaves from late-season, suspected glyphosate escapees. Provided are the Ct values for ALS and EPSPS targets, the differences of these two values (∆CT), the absolute difference, and the raw estimates of EPSPS gene copies (2∆CT). |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Species | Sample Name | ALS (Ct) | EPSPS (Ct) | ∆CT (intra-assay) | Abs (∆CT) | EPSPS Copies (2∆CT) |
| Monoecious Pigweed | 20-NP-33o | 24.303 | 25.023 | 0.721 | 0.7205 | 1.648 |
|  | 20-NP-33y | 24.132 | 24.923 | 0.791 | 0.7907 | 1.730 |
| Waterhemp | 20-NP-34o | 25.074 | 25.623 | 0.549 | 0.5491 | 1.463 |
|  | 20-NP-34y | 24.796 | 25.329 | 0.533 | 0.533 | 1.447 |
| Waterhemp | 20-NP-35o | 24.618 | 23.352 | -1.266 | 1.2657 | 2.404 |
|  | 20-NP-35y | 25.442 | 24.175 | -1.267 | 1.267 | 2.407 |
| Waterhemp | 20-NP-36o | 23.884 | 23.816 | -0.068 | 0.0681 | 1.048 |
|  | 20-NP-36y | 23.952 | 23.907 | -0.045 | 0.0451 | 1.032 |

# SECTION 9: 	SANGER SEQUENCING

Objective:

To confirm that PCR products produced by EPSPS-Copy assay from various samples aligned with the targeted region with the EPSPS and ALS genes.

Procedures:

Briefly, sample were amplified at each gene separately and PCR products were separated on a 1% agarose gel at 145V for 40 min. The corresponding 118bp (ALS) and 194bp (EPSPS) amplicons were excised from the gel, purified from the agarose and DNA binding dyes and eluted in a final volume of 20µl using a gel clean-up kit following manufacturer instructions (Wizard SV Gel and PCR kits). The eluted PCR products were sent to the UCCCC DNA Sequencing Facility (Chicago, IL) along with each primer (at 4µM final concentration) for Sanger sequencing in both directions using the corresponding forward and reverse primers from the assay. The amplified products came from plate 080420_1RT with corresponding gels 011121_1G. All four samples were selected for sequencing because they represented the different pigweed species (Table 9.1).

Results & Conclusions:

The initial agarose gel results displayed the expected banding pattern with ALS showing smaller product when compared to EPSPS gene (Fig 9.1). The Palmer amaranth samples (19-0055-4) is known to have a high EPSPS:ALS ratio (~88 copies), compared to the other samples that have a lower ratio (~2-6 copies).

Figure 9.1. Image of the 1% agarose gel for EPSPS and ALS targets. The relative size of the PCR products can be estimated by comparing the amplicons to the ladder (lane 1) and the corresponding base pairs to the left of the gel. NTC is showed in lane 10.

Sequencing analyses had high resolution for both ALS and EPSPS targets, which allowed NAGC to obtain the locations where the primers and probe anneal to each gene target. Sequence alignments for ALS and EPSPS sequences showed polymorphic sites within the amplicons, most of which are outside the primers and probes annealing locations. The polymorphisms are not surprising given that these samples represent the different species of pigweeds. Since these SNPs are within the exon or coding region for each gene, they are expected to be synonymous mutations. That is, these SNPs generally are silent mutations that do not change the amino acid sequence at that location. For the EPSPS sequences, there were two SNPS within the samples that are synonymous substitutions: 47th (A to G) and 53rd (C to A). The 53rd nucleotide position may be of interest as it results in both an amino acid substitution (A to D) and a potential change in the localized properties (hydrophobic to acidic, charged). For ALS, there was one SNP at 88th bp (C to T) that resulted in an amino acid change from Leucine (L) to Phenylalanine (F). Both the L and F are hydrophobic so the structure of the protein may not change due to this non-synonymous substitution.

The sequencing analyses confirmed that the EPSPS-Copy assay is amplifying the appropriate targets. The EPSPS sequences matched 100% with at least one sequence found within the NCBI database. In contrast, the database had a limited number of monoecious pigweed accessions or only had partial sequence data. In the cases where partial sequence information was available, the samples matched with sequences from the same, expected species. Overall, this demonstrates the universality of the EPSPS-Copy assay to detect for changes in EPSPS copy numbers as it relates to ALS in all pigweed species. The specific mechanisms for elevated EPSPS copy number variation (gene duplication) in the dioecious species may not be present in monecious pigweeds. Regardless, the assay can help further explore EPSPS in these species.

| Table 9.1. Summary of the samples selected for sequencing and the resulting match found in NCBI for the sequences analyzed using BLAST. For matches the species with 100% identical sites is indicated in the parentheses. Sequences are from submission 011321_NAGC. |  |  |  |
| --- | --- | --- | --- |
| Sample Name | Species | EPSPS Accession Match | ALS Accession Match |
| 19-0041-1 | A. retroflexus (RR) | KT833341 (RR) | KY781916 (PA; RR only partial in NCBI) |
| 19-0055-4 | A. palmeri (PA) | KC169785 (PA) | KM438057 (PA) |
| 19-0078-2 | A. tuberculatus (WH) | FJ869881 (WH) | EF157818 (WH) |
| 19-0085-3 | A. powellii (PO) | KM217581 (PO) | KM438056 (PA; PO only partial in NCBI) |

# SECTION 10: 	STANDARD SPECIMEN TESTING

Objective:

Previous work on the Amaranth Species ID assays created an archived client sample set to use to assess the EPSPS-Copy assay from case-like samples.

Procedures:

A selection of archived DNA samples from leaf samples were examined. All samples were previously extracted using established laboratory protocols at the time with some extractions using ChargeSwitch and others using the current Maxwell protocols. Specifically we tested archived extracts of Palmer amaranth (n = 26), waterhemp (n = 29), Powell amaranth (n = 4), and redroot pigweed (n = 4) that were collected in 2019 and 2020 from fields across Minnesota and North Dakota.

Results & Conclusions:

In total, 25% (14/55) of the dioecious pigweed samples had elevated EPSPS gene copies, the mechanism that confers resistance to glyphosate (Fig. 10.1). Interestingly, waterhemp samples showed elevated EPSPS copies just below the resistance threshold, signifying EPSPS copy number variation within this group of samples. There was no evidence of elevated EPSPS copies in the Powell amaranth or redroot pigweed samples, but it was a limited dataset.

Figure 10.1. EPSPS copy number variation in a subset of MN and ND pigweed samples submitted to NAGC in 2019 and 2020. The dotted line indicates the threshold of relative EPSPS gene copies associated with glyphosate resistance. The larger numbers of EPSPS gene copies in Palmer amaranth samples were expected due to the unique duplication mechanism (eccDNA, described in Section 8) that results in more copies, compared to waterhemp samples.

# SECTION 11: 	MEASUREMENT OF UNCERTAINTY

Objective:

For quantitative assays, there are several steps within the procedure that can introduce error into reported results. Thus, results from test samples are best estimates of the exact or true value, it is important to include calculations that account for these errors and assess repeatability (duplicates) and reproducibility (independent analysts) of results under various conditions. Here, these measures of uncertainty (MU) were calculated for the EPSPS-Copy assay so that a range of quantities may be included with the best estimate of quantities for reported samples.

Procedures:

The MU was calculated for each target of the EPSPS-Copy assay using equations 1 – 5 and procedures outlined in SOP 3.1 Measurement of Uncertainty v3.0. The dataset used in the uncertainty measurement came from various samples that spanned across the validation and different analysts. Due to the inclusion of multiple steps within the protocol, both within-analyst and between-analyst MU were included in the calculation.

Results & Conclusions:

In total, 29 (ALS) and 9 (CPS) samples were ran in duplicate for the final MU calculation of the EPSPS-Copy assay (Table 11.1). The ALS replicates were less variable than the EPSPS replicates as indicated by the RSD values. The largest disparity was two Cts was between synthetic DNA dilution replicates for EPSPS.

| Table 11.1: Summary of MU calculations for the EPSPS-Copy assay. RT plates and specific Ct values for all replicates are documented within the MU Worksheet – EPSPS-Copy folder on the NAGC drive. |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Target | No. of Analysts | No. of Replicates | Average Ct | ΣDi | SR´ | RSD |
| EPSPS (w/ALS) | 2 | 29 | 27.473 | 19.257 | 0.321 | 1.168 |
| ALS | 2 | 29 | 28.918 | 4.767 | 0.104 | 0.358 |
| EPSPS (w/CPS) | 2 | 9 | 24.289 | 0.039 | 0.002 | 0.009 |
| CPS | 2 | 9 | 26.196 | 0.079 | 0.004 | 0.017 |

Combined Uncertainty Measures Ct Values

By using Equation 5 of SOP 3.1 (v2.2), the combined relative standard uncertainty (RSU) for the assay was:

RSU = √ (1.168)2 + (0.358)2 + (0.009)2 + (0.017)2 = 1.222%

Therefore, the expanded uncertainty (U) when using a coverage factor of 2 gives a level of confidence of approximately 95%:

U = 1.222% * 2 = 2.444%

No diagnostic assay is 100% efficient and reported values are best estimates of the true value. Thus, it is important to include measures of uncertainty (MU) with quantitative results, especially for tests that report pathogen loads of samples. Due to the lack of certified reference material and lack of other laboratories performing the EPSPS-Copy assay inter-laboratory proficiency comparisons, MU was calculated from relative intra-laboratory reproducibility measures using Ct replicates from the standard curve calculations and PAC replicates. The presented MU calculations are well within the range of those previously collected for other quantitative tests at NAGC. If clients request more information on uncertainty measures, the following explanation can be provided for the EPSPS-Copy assay:

“For all quantitative assays, the reported result for each sample is the best estimate of the exact or true value. Thus, results are reported with a range of values that are likely to include the true value. This range of values can be calculated by multiplying the best estimate by a relative measure of uncertainty (%). This calculated value is then added or subtracted from the best estimate to generate the range of values. The expanded relative measure of uncertainty is 2.44% for the EPSPS-Copy assay. Thus, the true value of a particular sample will likely be found within ±2.44% of the best estimate (EPSPS copy number) reported within a confidence interval of approximately 95%.”

The EPSPS-Copy assay differs from the typical quantitative assays for pathogen diagnostics at NAGC because it does not require a standard curve for quantification. Instead, this assay directly compares the Ct values between two targets for relative quantification of EPSPS within each sample. Nevertheless, a standard curve should still be ran following dye calibrations to monitor the PCR efficiencies of each target. Using 1:5 dilutions of a real samples provided the best method to calculate PCR efficiencies in this validation and is recommended for future quality assurance tests.

# SECTION 12: 	GREEN HOUSE COMPARISON

Objective:

A singleplex version of the EPSPS-Copy assay has been in use since Gaines et al. (2010) for a variety of statewide surveys to estimate the frequency of glyphosate resistance genotypes in the field. As an ISO-accredited laboratory, NAGC invites any opportunity to compare an assay with another detection or diagnostic method. For glyphosate resistance, it is important to investigate whether genotypes match the expected phenotype. It is known from previous research that the EPSPS-Copy assay will not explain 100% of the resistant pigweeds in the field as there is an additional known target-site mutation (SNP; P106S) as well as a non-target-sit mechanism that results in reduced glyphosate translocation (Tranel 2020). Despite these different mechanisms that confer glyphosate resistance, it is still important to compare greenhouse and EPSPS-Copy assays as a proof-of-concept for the practical application of genotyping as a method to detect herbicide resistance in pigweed.

Procedures:

Our collaborator, Dr. Joe Ikley (NDSU weed extension specialist), had a greenhouse herbicide trial of Palmer amaranth seedlings near the completion of this validation. The seeds were collected from sunflower screenings so seed origin was unknown and could have been diverse, representing plants from multiple field or states. Prior to spraying seedlings in the greenhouse with glyphosate, Dr. Ikley collected a single leaf sample from 112 seedlings for NAGC to test with the EPSPS-Copy assay. Importantly, the results of the glyphosate application were blinded to NAGC and only released once NAGC sent the results from the genotyping test. A research-use only report was provided to Dr. Ikley (Case No 20-0177), where results were reported as relative EPSPS gene copies

Results & Conclusions:

When using the relative measure of four (4) EPSPS gene copies as a guideline for potential glyphosate resistance, there was a 90% (101/112) agreement between genotyping and phenotyping during this blinded study (Table 12.1). There appears to be a range of recommendations for the minimum number of EPSPS copies necessary to confer glyphosate resistance, which has been based on comparative results between genotyping and the shikmate assay. To date, the minimum has been 4 copies (∆Ct = 2), but others have suggested 10 (∆Ct = 3.3) copies or more (Dillon et al. 2017; Giacomoni et al. 2019). There has also been some discussion that EPSPS copy number and resistance depends on the species, specifically that resistance in Palmer amaranth should result in many more copies due to eccDNA mechanisms described in Section 8. In contrast, waterhemp can be resistant at much lower copies than Palmer amaranth. Further research has suggested that there may be a dosage response where more EPSPS copies provides greater glyphosate protection in pigweeds (Giacomoni et al. 2019). In our small study, adjusting the threshold to 4 copies (∆Ct = 2) provided a better match for the genotype-phenotype comparison.

In the samples where genotypes and phenotypes did not match (disagreement group), two glyphosate resistant plants had an estimate of 3 EPSPS copies. To be conservative, these were denoted as a susceptible genotype (∆Ct = 1.4 & 1.6). There were also five plants that were glyphosate resistant, containing 1 or 2 EPSPS copies per haploid genome. These low copy but resistant plants could have additional mechanisms to explain glyphosate resistance. One study estimated a false negative or mismatch rate of 18%, a rate at which we are below in our small study here (10%) (Chatham et al. 2015). Interestingly, there were 3 disagreements where NAGC found elevated levels of EPSPS (8, 10, and 15) in glyphosate susceptible plants. According to Dr. Ikley, one of these may have been an accidental labeling mistake in the greenhouse for a sample. Since no studies published any false positive rates, the mechanisms of susceptibility in the face of elevated EPSPS is unknown. Outside of sample mislabeling, other environmental factors may have contributed to their untimely demise.

| Table 12.1. Summary of the blinded study for glyphosate resistance in Palmer amaranth seedlings from sunflower screens. Raw data can be found in the Greenhouse Study worksheet within the 1.0 Complete Summary of EPSPS-Copy Data workbook in the validation folder. |  |  |  |  |
| --- | --- | --- | --- | --- |
|  |  | Phenotype |  |  |
|  |  | Resistant | Susceptible | Total |
| Genotype | Resistant | 39 | 3 | 42 |
|  | Susceptible | 7 | 63 | 70 |
|  | Total | 46 | 66 | 112 |

# SECTION 13: CPS – ALS COMPARATIVE STUDY TO CHANGE REFERENCE GENE TO CPS

Objectives:

Since ALS gene is likely under strong selective pressure from ALS-inhibiting herbicides, there were criticisms within the literature about using the ALS gene as a reference marker for the EPSPS-Copy assay. An alternative marker, the CPS gene, is considered a single-copy gene (Ma et al. 2013) that is not directly impacted by herbicide or other known selective pressures. This section of validation documents the primer-probe bioinformatics and subsequent qPCR results that support the use of the CPS instead of ALS for the reference gene within the EPSPS-Copy assay

Procedures:

Earlier in the validation, there was no public sequence information available to examine the PCR product of Ma et al. (2013)’s EPSPS assay that used the CPS gene as a selectively neutral reference gene. The original assay used SYBR chemistry, but NAGC needed to develop a probe so it can be co-amplified as a duplex assay with EPSPS. Sanger Sequencing the PCR product using the Ma et al. (2013) method was not possible because the product is too short to get adequate resolution for nucleotide analysis. Fortunately, genomic sequences from waterhemp, smooth pigweed, and Palmer amaranth became publicly available within the database CoGe (https://genomevolution.org/coge) through the recent work of Montogmery et al. (2020). With this new genomic information, novel primers and probes could be developed within the CPS gene suggested by the original Ma et al. (2013) assay.

Once primers and probes were developed in silico, three combinations of primers and probes were tested in the same reaction as the EPSPS target (duplex; Table 13.1). Templates for these master mix trials included a series of 1:5 dilutions of a waterhemp extract (20-NP-39) as well as a known sample (19-0055-4.1) with elevated EPSPS copies compared to the ALS (∆Ct = 5) (Experiment 1). Next, specificity of the CPS target was examined with other laboratory-related PACs and various crop extracts. Lastly, a large sample set of greenhouse screenings (described in Section 12), were re-analyzed using the EPSPS-CPS duplex to compare to the results of the original ESPPS-ALS estimates. The collective analyses of Section 13 provided the recommendation to use CPS as given below.

| Table 13.1. Primer and Probe combinations used in the initial CPS assay trial. Provided are the master mix (MM) designations, names of oligos, and the respective sequences. Ma et al. (2013)’s CPS assay was identified during the bioinformatics and their primers were tested with primers and probes developed at NAGC (see Source column). |  |  |  |  |
| --- | --- | --- | --- | --- |
| Master Mix (MM) | Primer F Name: Sequence (5′-3′) | Probe Name [HEX]: Sequence (5′-3′) | Primer R Name: Sequence (5′-3′) | Source |
| 1 | CPS-F2 | CPS-Pr2 | CPS-R2 | NAGC (F/P/R) |
| 2 | CPS-F3 | CPS-Pr1 | CPS-R1 | NAGC (F/P) Ma et al. 2013 (R) |
| 3 | CPS-F1 | CPS-Pr2 | CPS-R1 | NAGC (P) Ma et al. 2013 (F/R) |

## Experiment 1: Primer and Probe Combination & PCR Efficiency Comparisons

Procedure:

Three master mix combinations in Table 13.1 with different forward and reverse primers as well as two probes were tested at the reagent concentrations and thermal cycling conditions that were optimized for the EPSPS-ALS assay. Titrating the primers and probes for the CPS target was not investigated because one target should not be preferentially amplified. By using a series of 1:5 dilutions from a known waterhemp sample, PCR efficiencies could be used to find a CPS assay that closely matched the efficiency as the EPSPS assay.

Results & Conclusion:

Of the three master mixes that were examined, Master Mix 2 (MM2) provided the closest PCR efficiencies with EPSPS target and provided the highest fluorescence as a duplex assay (Table 13.2). The EPSPS-CPS assay with MM2 had similar slopes as the EPSPS-ALS assay, but also had improved fluorescence. The closeness of the PCR output between ALS and CPS within the dilution series provides evidence that CPS is an appropriate replacement marker.

| Table 13.2. Summary of the PCR efficiencies calculated for the three master mixes (MM) ran as a duplex with the EPSPS assay. For comparisons, the EPSPS-ALS assay efficiencies (in green) are from a separate run (Plate: 021221_2RT). Data for all MMs were calculated from plate 032321_3RT. MM2 in blue had similar PCR profiles with the EPSPS-ALS assay. |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Assay Master Mix | Target | Slope | Y-int | R2 | % Eff | Max Fluorescence |
| MM1 CPS-F2/R2/Pr2 | EPSPS | -2.998 | 21.431 | 0.997 | 115.526 | 3.2 million |
|  | CPS | -3.526 | 23.960 | 0.985 | 92.135 | < 1 million |
| MM2 CPS-F3/R1/Pr1 | EPSPS | -3.438 | 21.117 | 1.0 | 95.366 | 3.5 million |
|  | CPS | -3.392 | 22.661 | 1.0 | 97.137 | 2.7 million |
|  | EPSPS | -3.440 | 22.653 | 0.999 | 95.309 | 3.3 million |
|  | ALS | -3.373 | 24.323 | 0.999 | 97.916 | 2.2 million |
| MM3 CPS-F1/R1/Pr2 | EPSPS | -3.447 | 21.090 | 1.0 | 95.021 | 3.2 million |
|  | CPS | -3.343 | 24.068 | 0.999 | 99.116 | 1.5 million |

## Experiment 2: Specificity of CPS gene assay

Procedure:

The purpose of the specificity analysis was to evaluate the primers and probe of MM2, checking for cross-reaction to non-specific targets. DNA from laboratory, agriculturally, and environmentally-related samples were previously extracted using established laboratory protocols. These samples were screened using the EPSPS-Copy assay to verify cross-reactivity with other DNA samples that occur in the laboratory. In addition, several other amaranth samples were included to compare EPSPS-CPS assay to the previous EPSPS-ALS assay. The samples included single seeds from the five common pigweed species as well as a sample with an elevated copy number of EPSPS. The cycling parameters and reaction conditions remained the same, with the substitution of CPS (MM2; Table 13.1) for ALS in the EPSPS-Copy assay.

Results & Conclusions:

The specificity analysis revealed one potential non-target amplification for the EPSPS target (Table 13.3), which was discovered earlier in the validation (see Section 6). No other plant or tested organism, besides pigweed species, produced both CPS and EPSPS signals. The CPS gene assay appears to be as specific as the ALS gene assay.

| Table 13.3 Specificity results (Ct values) for EPSPS and CPS targets. Green shading indicates the potential cross-reaction with a Fusarium species. Data analyzed from plate 032421_1RT. |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Sample Name | EPSPS | CPS | ∆Ct |  | Sample Name | EPSPS | CPS | ∆Ct |
| 20-NP-39 | 22.385 | 24.200 | 1.814 |  | Soybean | Und | Und | - |
| Amaranth PAC | 22.226 | 24.992 | 2.766 |  | LAB Ctrl | Und | Und | - |
| Palmer amaranth | 20.929 | 24.146 | 3.217 |  | DCV PAC | Und | Und | - |
| Redroot pigweed | 25.621 | 25.030 | -0.591 |  | Neg Bee | Und | Und | - |
| F. graminerium | 38.309 | Und | - |  | Sugarbeet | Und | Und | - |
| ABL2 PAC | Und | Und | - |  | Corn | Und | Und | - |
| GLX PAC | Und | Und | - |  | NCLB PAC | Und | Und | - |
| Potting Soil (16-NP-318) | Und | Und | - |  | Barley | Und | Und | - |
| NTC | Und | Und | - |  | Lentil | Und | Und | - |
| CLRDV PAC | Und | Und | - |  | Wheat DNA | Und | Und | - |
| AES PAC | Und | Und | - |  | 19-SCN-19 | Und | Und | - |
| F. subglutinans | Und | Und | - |  | PS PAC | Und | Und | - |
| Salmon sperm DNA | Und | Und | - |  | Canola | Und | Und | - |
| Sandy Loam (16-NP-108) | Und | Und | - |  | Pea | Und | Und | - |
| DIP PAC | Und | Und | - |  | Elk Muscle | Und | Und | - |
| KNN PAC | Und | Und | - |  | 19-SNC-4 | Und | Und | - |
| F. verticillioides | Und | Und | - |  | Puccinia PAC | Und | Und | - |
| Spinach | Und | Und | - |  | Chickpea | Und | Und | - |
| GW PAC | Und | Und | - |  | Rice | Und | Und | - |
| ISL1 PAC | Und | Und | - |  | Clay Soil (16-NP-317) | Und | Und | - |
| Human DNA | Und | Und | - |  |  |  |  |  |

The ∆Cts were consistent between the single seed samples and the leaf with elevated EPSPS copies, regardless of the reference gene (ALS vs CPS) used in the assay (Table 13.4). These data indicated that the newly developed EPSPS-CPS assay can be used to examine EPSPS copies in a variety of amaranth species and sample sources (leaf and seed). The next step was to compare a larger dataset to examine any differences between the two EPSPS-Copy assays with client samples.

| Table 13.4. Comparison of the assays with different reference markers on single seed samples from various species and a known amaranth sample with elevated EPSPS copies (19-0055-4.1). The ΔCt values are absolute values of ALSCt-EPSPSCt or CPSCt-EPSPSCt. |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Sample Name | Plate: 021221_2RT |  |  |  | Plates: 032421_1RT & 032321_3RT |  |  |
|  | EPSPS | ALS | ΔCt |  | EPSPS | CPS | ΔCt |
| Redroot seed 120419LP | 29.412 | 28.056 | 1.356 |  | 29.189 | 28.373 | 0.816 |
| Powell seed 120419LP | 29.127 | 27.987 | 1.140 |  | 28.386 | 27.596 | 0.790 |
| Palmer seed 120419LP | 28.895 | 27.630 | 1.265 |  | 28.538 | 27.804 | 0.734 |
| Waterhemp seed 120419LP | 31.158 | 30.455 | 0.704 |  | 30.613 | 29.944 | 0.670 |
| Smooth seed 120419LP | 29.826 | 28.616 | 1.210 |  | 29.580 | 28.929 | 0.652 |
| 19-0055-4.1 (1:50) | 23.980 | 29.033 | 5.083 |  | 22.590 | 27.657 | 5.067 |

## Experiment 3: Greenhouse Standard Specimen Testing

Procedure:

Our collaborator, Dr. Joe Ikley, had a greenhouse herbicide trial of Palmer amaranth seedlings near the completion of this validation (see Section 12). Briefly, seeds were collected from sunflower screenings so seed origins were unknown and could have been diverse, representing plants from multiple field or states. Prior to spraying seedlings in the greenhouse with glyphosate, Dr. Ikley collected a single leaf sample from 112 seedlings (Case No. 20-0177) and 92 seedlings (Case No. 21-0019) for NAGC to test with the EPSPS-Copy assay. Both cases were originally tested and reported using ALS as the reference gene. We re-analyzed these samples (n = 204) using the newly adapted EPSPS-Copy assay using CPS as the reference gene. Samples were tested just once using the EPSPS-ALS and EPSPS-CPS assays.

Results & Conclusions:

Overall, the results were relatively congruent between the EPSPS-ALS and EPSPS-CPS assays. The side-by-side comparison of the 204 samples can be viewed in worksheet ALS vs CPS Target Comparison within the workbook CPS Analysis Summary of the validation. Absolute differences between the Ct values of the targets and ∆∆Ct calculations are shown in Table 13.5. The Ct values remained similar for the EPSPS target within tested samples, regardless of the assay being performed. The maximum difference of the EPSPS target was at or below 0.5 Cts for any sample tested with both duplex assays. As expected, there was a larger difference in Ct values between the two single-copy reference genes, with a maximum difference of just 1.3 Cts. The maximum change of 0.89 in the ∆Ct between the assays, indicates that replicates may differ by approximately 1 power when using the 2∆Ct method. While at large Ct differences (i.e., elevated EPSPS copies) the estimates can differ substantially (24 = 16 vs 25 = 32), both assays show congruency when calling low EPSPS copies.

Interestingly, there appears to be an inflection point in this dataset that falls along the 4-copy threshold established for glyphosate resistance.

| Table 13.5. Descriptive statistics when comparing the two EPSPS-Copy assays using different reference (i.e., single-copy) genes on Palmer amaranth leaf samples from cases 20-0177 and 21-0019 (n = 204). These summary statistics were calculated by calculating the absolute values of the Ct differences for the EPSPS target between both duplexes (EPSPS-EPSPS) and between the reference genes (CPS-ALS). The ΔΔCt column reflects the change (Δ) in the ΔCts when subtracting the ΔCts of EPSPS-ALS assay from the ΔCts of the EPSPS-CPS assay. |  |  |  |
| --- | --- | --- | --- |
| Statistic | EPSPSCt - EPSPSCt | CPSCt - ALSCt | ΔΔCt |
| Average | 0.10 | 0.52 | 0.48 |
| Standard Deviation | 0.08 | 0.21 | 0.19 |
| Maximum | 0.50 | 1.30 | 0.89 |
| Minimum | 0.00 | 0.01 | 0.05 |

Figure 13.1. Graphic summary of the comparison between the two EPSPS-Copy assays. EPSPS copies with true rounding are plotted on the y-axis on Plot A. The x-axis indicates the sample order, a number that was assigned to align the assay results from the same sample but sorted by the ΔCt (high to low) provided by the EPSPS-CPS assay. This sorting is why the EPSPS-CPS plotted data appears more linear (and less noisy), compared to the EPSPS-ALS plotted data. While difficult to see individually, each sample has two dots (blue and red) that are stacked vertically on the graph. The green line indicates the location of the 4 copy threshold for EPSPS. Note that below the threshold, both EPSPS-Copy assays stay consistent on their qualitative result: low EPSPS copies. Plot B shows the data at a higher level with the ΔCt plotted on the y-axis. The x-axis labeling and sorting is identical to Plot A, but shows an inflection point that is obvious in both assays. This gap spans about 0.5 ΔCt with nearest ΔCt at 1.97 and 1.49.

Conclusions:

The collective results from this section provide enough evidence to support the switch from ALS to CPS as the reference gene to calculate relative EPSPS copies in amaranth plants. The CPS target appears to be just as specific as the ALS target in amaranth species and closely matches the PCR efficiencies with the EPSPS target. Finally, a large dataset of samples were re-analyzed and the qualitative results closely matched, regardless of the reference or single-copy gene used.

# SECTION 14: 	NEGATIVE CONTROL CONTAMINATION ASSESSMENT

Objective:

Determine the contamination rate and level, if any, present in the negative control samples (reagent controls and no template control samples).

Procedures:

In order to assess a contamination rate and/or level, each RC and NTC from every EPSPS-Copy assay during the validation was analyzed, keeping track of any assay that had a positive Tm profile or Ct value in the negative controls. Any negative control sample producing a possible signal was individually analyzed to determine if a contamination event actually occurred.

Results & Conclusions:

The results were tabulated within a workbook, labeled Negative Control Assessment worksheet within the located within the 1.0 Complete Summary of EPSPS-Copy Data workbook in the validation folder. The CPS data was also incorporated into the Negative Control Assessment worksheet.

In summary, contamination data was obtained from analyzing 25 probe-based plates and 2 HRM plates, which totaled 54 negative controls (3 RCs and 51 NTCs). The minimal RCs within this validation was because many archived DNA samples were used in this study. Therefore RCs were only included when samples being analyzed were extracted immediately prior to the PCR assay and/or for the first time of analysis; however, an NTC reaction was included on every PCR assay.

The contamination rate for the EPSPS-Copy assay was 3.7% (2/54), having two independent contamination events occurring during the assay validation (54 assays). The cross-contaminated sample happened to be two RC samples from the greenhouse study (Section 12). Only the ALS and CPS target was present at a high Ct (37 & 38), suggesting a potential contamination event in plate 120920_2RT and 032921_2RT. Re-amplifying the same RC showed the presence of the ALS target (021221_2RT), indicating contamination during the extraction process on the Biomek.

# SECTION 15: 	PAC PREPARATION

Amaranth leaf tissue (preferred):

Bulk NA extract from any pigweed species could be used as a PAC for the EPSPS-Copy assay. The current Amaranth Species ID PACs (lot: 042919) consist of a mixture of the three target (all five species examined in the Amaranth Species ID validation). To minimize the need for an extra PAC, the Amaranth Species ID PAC can be used as long as it consists of amaranth species DNA from leaf material using current SOP extraction methods. While seed extract could be used, leaf material would be preferred. Only a single PAC is needed to run the assay as both targets (EPSPS and CPS) are within all amaranth species NAGC has tested.

Dye-Recalibrations – It is recommend to examine the PCR efficiencies for each target after recalibrations of the QuantStudio Flex Real-Time PCR Systems. For these measurements, it is best to perform six 1:5 serial dilutions of the Amaranth PAC or other leaf extract. PCR efficiencies can be determined in the same fashion as standard curve analyses.

gBlock synDNA (alternative option):

As a secondary option for EPSPS-Copy assay PACs, a mixture of synthetic DNA stocks or gBlock gene fragments purchased from IDT can be used. Obtain the synthetic EPSPS and ALS gene fragments of 244 bp and 168 bp, respectively (see Appendix C). Resuspend the gBlocks in TE-4 buffer to a final concentration of 10ng/µl, following instructions listed on the specification sheet. Briefly, add TE-4 buffer to the gBlock tubes and vortex vigorously for 1 minute. Spin the tubes briefly to collect the droplets to the bottom of the tube before incubating at 50°C for 20 minutes. Prepare 10-fold serial dilutions using MGW until an appropriate and equal Ct value for each target was achieved. Mix two target dilution tubes together. Verify using the EPSPS-Copy assay prior to PAC usage.

As with all other validations, the PAC needs to be analyzed in duplicate on 3 separate plates or have 6 replicates on the same plate in order to determine the Ct value range that falls within +/- 3SD of the average. The Ct value range can be used for estimating PCR efficiency for assay to assay variation and quantity extrapolation.

# SECTION 16: 	REFERENCES

Chatham, L. A., Bradley, K. W., Kruger, G. R., Martin, J. R., Owen, M. D., Peterson, D. E., & Tranel, P. J. (2015). A multistate study of the association between glyphosate resistance and EPSPS gene amplification in waterhemp (Amaranthus tuberculatus). Weed Science, 63(3), 569-577.

Dillon, A., Varanasi, V. K., Danilova, T. V., Koo, D. H., Nakka, S., Peterson, D. E. et al. (2017). Physical mapping of amplified copies of the 5-enolpyruvylshikimate-3-phosphate synthase gene in glyphosate-resistant Amaranthus tuberculatus. Plant physiology, 173(2), 1226-1234.

Gaines, T.A., Shaner, D.L., Ward, S.M., Leach, J.E., Preston, C. and Westra, P. (2011) Mechanism of resistance of evolved glyphosate-resistant Palmer Amaranth (Amaranthus palmeri). Journal of Agricultural and Food Chemistry 59: 5886-5889.

Giacomini, D., Westra, P., & Ward, S. M. (2014). Impact of genetic background in fitness cost studies: an example from glyphosate-resistant Palmer amaranth. Weed Science, 62(1), 29-37.

Giacomini, D. A., Westra, P., & Ward, S. M. (2019). Variable inheritance of amplified EPSPS gene copies in glyphosate-resistant Palmer amaranth (Amaranthus palmeri). Weed Science, 67(2), 176-182.

Gaines, T.A., Zhang, W., Wang, D., Bukun, B., Chisholm, S.T., Shaner, D.L., Nissen, S.J., Patzoldt, W.L., Tranel, P.J., Culpepper, A.S., Grey, T.L., Webster, T.M., Vencill, W.K., Sammons, R.D., Jiang, J., Preston, C., Leach, J.E. and Westra, P. (2010) Gene amplification confers glyphosate resistance in Amaranthus palmeri. PNAS 107: 1029-1034.

Koo DH, Molin WT, Saski CA, Jiang J, Putta K, Jugulam M, Friebe B, Gill BS (2018) Extra-chromosomal circular DNA (eccDNA) based amplification and transmission of herbicide resistance in crop weed Amaranthus palmeri. PNAS 115:3332–3337

Ma, R., Kaundun, S.S., Tranel, P.J., Riggins, C.W., McGinness, D.L., Hager, A.G., Hawkes, T., McIndoe, E. and Riechers, D.E. (2013) Distinct detoxification mechanisms confer resistance to mesotrione and atrazine in a population of waterhemp. Plant Physiology 163: 363-377.

Montgomery, J. S., Giacomini, D., Waithaka, B., Lanz, C., Murphy, B. P., Campe, R., ... & Tranel, P. J. (2020). Draft Genomes of Amaranthus tuberculatus, Amaranthus hybridus, and Amaranthus palmeri. Genome biology and evolution, 12(11), 1988-1993.

Murphy, B. P., & Tranel, P. J. (2019). Target-site mutations conferring herbicide resistance. Plants, 8(10), 382.

Tranel, P. J. (2021). Herbicide resistance in Amaranthus tuberculatus. Pest Management Science, 77(1), 43-54.